## Slide 1
Practical Multicolor
Flow Cytometry
Enrico Lugli, PhD
Laboratory of Translational Immunology
Humanitas Flow Cytometry Core
Istituto Clinico Humanitas, Milan, Italy

## Slide 2
Achille Anselmo
Technologist, specialist
Federico Colombo
Cell sorting specialist
The Humanitas Flow Cytometry Core
Enrico Lugli
Head
Principal Investigator

## Slide 3
How to avoid artifacts in flow cytometry
Reagent titration: how to achieve the best separation over background
Multicolor (>1 parameter) compensation
Spreading error in multicolor FACS
Presentation of high quality data for publication

## Slide 4
Abs: not just the manufacturer’s instructions
Fresh PBMC stained with 7 two-fold serial dilutions of BV421 anti-CD14 (uL of Ab in 100 uL total volume for 20’ at RT)
Titer:

## Slide 5
Abs: not just the manufacturer’s instructions
AIM: Choose the titer at which the best separation between the positive and the negative is achieved without increase in the background.
5
2.5
1.25
0.6
0.3
0.15
0.07
Titer

## Slide 6
Abs: not just the manufacturer’s instructions
AIM: Choose the titer at which the best separation between the positive and the negative is achieved without increase in the background.
✓
5
2.5
1.25
0.6
0.3
0.15
0.07
Titer
SNR: Signal-to-noise ratio
MFI: median fluorescence intensity
SNR: MFI (pos) / MFI (neg)
Titer
Chart: 
| 5 | 2.5 | 1.25 | 0.63 | 0.31 | 0.15 | 7.0000000000000007E-2
SNR | 219 | 230 | 247 | 283 | 265 | 264 | 249
SNR
✓

## Slide 7
Abs: not just the manufacturer’s instructions
Saturating titer: titer at which all receptors are saturated (no increase in positive signal despite the increase of the Ab concentration)
Separating titer: titer at which separation between the positive and the negative is obtained

## Slide 8
Titering antibodies: procedure
Start from a given amount of the antibody (generally the concentration suggested by the manufacturer or twice as much) in 100 uL total volume
prepare 8 two-fold serial dilutions, plus unstained
Add cells to each tube (number of cells IS IRRELEVANT – it can be 500K as 10 million)
Incubate at the desired condition (we do 20’ RT)
20 uL
10
5
2.5
1.25
0.6
0.3
0.15
Ab: 40 uL
Buffer: 60 uL
Final concentration:
Buffer: 50 uL
Transfer 50 uL from one tube to the other

## Slide 9
Careful with BV dyes

## Slide 10
Critical parameters affecting titers
TIME: the longer the incubation the lower the titer needed to reach saturation
TEMPERATURE: some Abs stain better at certain temperatures, e.g. chemokine receptors give a much higher fluorescence when stained at 37°C compared to RT or 4°C
BUFFER: intracellular staining (Fix/Perm buffers) can alter the optimized titer  titrate antibodies with the buffer used in the staining
Cell number DOES NOT affect the Ab titer. The Ab is always in vast excess to the number of molecules, unless cell number changes dramatically.
General rule: the optimized titer can be used to stain up to 10-15 x106 cells in 100 uL total volume

## Slide 11
Pro and cons of antibody titration
PRO:
Can save a lot of money (100 tests can become 400)
Better separation, less background
VERY important when performing multicolor analysis (can minimize “spreading” in other channels)
CONS:
None

## Slide 12
Optimized voltage settings in FACS experiments

## Slide 13
Voltage calibration
Historically, voltage setting has been determined empirically, i.e. voltage is increased (or decreased) to a value capable to show our positive cells
Intuitive when very bright signals are analyzed (e.g., CD8, CD14, GFP)
Not clear when dim antigens are investigated (chemokine receptors, phosphoepitopes)

## Slide 14
Voltage calibration
Voltage
Optimized
Increased background
Decreased separation
Optimized voltages to generate the best possible separation over background. Optimized for dim signals
If analyzing bright signals, turn down voltage to bring your cells on scale
Perfetto SP, Nat Protoc, 2012

## Slide 15
Fluorescence Compensation

## Slide 16
Compensation
Procedure through which the “unwanted” fluorescence from the primary channel (FITC) is removed from the secondary channel (PE)

## Slide 17
Compensation: just linear algebra
Primary detector
Secondary detector
ca. 13,000
ca. 1,100
Compensation:
1,100:13,000 = x:100
x = 1,100*100/13,000 = 8.46
LINEAR subtraction of the contaminating fluorescence signal from the primary detector
The same value will be subtracted from all the cells in the fluorescence range

## Slide 18
Is this correctly compensated?
Credit: Mario Roederer, VRC, NIH

## Slide 19
Aspects to consider
Compensation can be performed manually (at the machine) or automatically (via software)
Manual compensation: tedious, need to check carefully that positive and negative signals have to match
Automated compensation: no errors, if using the right controls
Manual compensation is impractical in a multicolor experiment (roughly 2n pairwise combinations)
Changing compensation in a pairwise combination affects distribution in other channels.

## Slide 20
Practical compensation
How do we correctly estimate the “spillover” of a given fluorochrome into other channels?
SINGLE-STAINED CONTROLS

## Slide 21
Compbeads for single stained controls
Plastic beads covered with species-specific anti-Fc antibodies (anti-mouse, -rat, -hamster, etc.)
Very bright
PROCEDURE:
Vortex beads
Aliquot 30 uL to each tube (including a negative control)
Add Ab (the same titer used in the staining)
Vortex
Incubate 15’ RT
Wash

## Slide 22
Minimizing compensation
Good 4 color combinations:
FITC – PE-Cy7 – PacBlue – APC
FITC – PE – PacBlue – APC
FITC – PE – PacBlue – APC-Cy7
Success of a panel does not depend on compensation. Panels work fine even when compensation is >100%
Spreading error between channels determines the success of the panel

## Slide 23
The 3 golden rules of compensation
Use a single-stained control that is AT LEAST as bright as the test staining
Use the same reagent that is being used in the staining (valid for tandems such as PE-Cy7, APC-Cy7, etc. – for non tandems any reagent is OK)
Have a paired-negative control with the SAME autofluorescence as the positive (cells with cells, beads with beads, etc.)
Everything else is not necessary!

## Slide 24
Experiment plan
A 5-color panel is designed to investigate CD3 FITC, CD4 APC, CD8 Cy5PE, CD19 PE and Aqua live/dead (to exclude dead cells)
Tubes:
  Unstained comp control
  Beads: CD3 FITC
  Beads: CD4 APC
  Beads: Cd8 Cy5PE
  Beads: CD19 PE
  Beads: Aqua (positive and negative)
  Full stained samples
  8-12. FMO controls (up to 5)

## Slide 25
Rule #1: at least as bright!
FITC IFN-g
Alexa 594 CCR7
Overcompensation of FITC into Alx594
Comp control
Test staining: IFNg
Comp control is not as bright as the test stain

## Slide 26
How to accomplish Rule #1?
Get a brighter comp (stain with more Ab; use a different Ab, binding beads at higher efficiency, if the Ab is not a tandem)
Decrease Ab titer in staining (not in this case, you will have loss of sensitivity)
Use single-stained cells

## Slide 27
Rule #2: different Abs conjugated to the same FL require different compensation
Compensation requirements can be different from lot to lot of the same Ab
Wrong comp
13.2% subtraction
PE-Cy5.5
Correct comp
3.4% subtraction
PE
Negative control
PE-Cy5.5
PE
Wrong
Correct

## Slide 28
Changing compensation affects distribution in other channels
FITC (l 525) and PerCP (l 670) are distant enough
Possibly wrong comp between PerCP (l 670) and PE-TxR (l 610)

## Slide 29
How to compensate in FACSDiva
Choose Experiment  Compensation setup  Create compensation controls
Run single-stained controls, including negative
Select positive and negative signal from each control
For Comp tubes containing both positive and negative, draw an additional gate to select the negative population
Choose Experiment  Compensation Setup  Calculate Compensation.
Enter a compensation setup name, click Link and Save, and click OK.

## Slide 30
Spreading error in the muticolor world

## Slide 31
Signal spreading
Compensated
Compensated with transformation
Uncompensated
1000
2000
-500
+500
What causes this spread? Photon counting statistics introduces a measurement error
Credit: Mario Roederer, VRC, NIH

## Slide 32
“Errors” in compensation
Properly compensated data may not appear rectilinear because of measurement errors (are more prominent in case of limited numbers of photons, Cy7PE or Cy7APC dyes, with low energy)
Measurement errors are directly proportional to the intensity (LOG relationship)
This effect on compensated data is unavoidable, and cannot be corrected with compensation (LINEAR correction)
Compensation does not introduce the error, it only REVEALS it because it brings the distribution to the low end of the log scale

## Slide 33
How do we overcome this problem?
Use different gates: cells positive in the primary detector will have a different threshold of positivity compared to negative cells
APC
Alexa 700

## Slide 34
Use different gates: cells positive in the primary detector will have a different threshold of positivity compared to negative cells
Not a problem if analyzing bright antigens in the secondary detector
APC
Alexa 700
100% detected
How do we overcome this problem?

## Slide 35
Use different gates: cells positive in the primary detector will have a different threshold of positivity compared to negative cells
Not a problem if analyzing bright antigens in the secondary detector
Loss of sensitivity if DIM antigen in the secondary detector. Decrease titer of spreading Ab to minimize spreading.
Try a different Ab combination
APC
Alexa 700
100%
detected
~50%
detected
How do we overcome this problem?

## Slide 36
Final considerations for panel development
Ideally not a problem if compensation is very high between fluors (need to be tested in any case)
Voltage amplification determines compensation values but not the final result (we are compensating fluorescence)
No need to set voltages such that the signal in the primary channel is brighter than that in the secondary channel
Spread of signals determines the success of multicolor panels
Primary detector
Secondary detector
Compensation: 100%
Compensation: <100%
Compensation: >100%

## Slide 37
Define markers to be used
Dead cell exclusion: Aqua Zombie live/dead discriminator
Lineage: CD3, CD4, CD8, CD45 (optional?  save one channel)
Memory differentiation: CD45RO, CCR7, CD95, CD27, CD57 (minimum set of markers to identify naïve and memory T cell subsets – Lugli, Nat Protoc, 2013)
Defining T cell exhaustion and activation: PD-1 (exhaustion), HLA-DR (activation), Granzyme B and Perforin (cytotoxicity; T cells may still be cytotoxic even though inhibited), T-bet and Eomes (effector functions)
Resident or circulating memory pattern: CD69

## Slide 38
Is this panel going to work?

## Slide 39
It is difficult to predict, especially the future

## Slide 40
Titrate every single reagent
Doesn’t only define correct concentration and if reagents are working, but also gives information to predict spreading error in a given channel  success of the multicolor panel
CD57 BV605
Concentration
Fluorescence
16
17
16
14
14
12
10
BV605
Cy5PE
Roederer, Cytometry, Cytometry. 2001 Nov 1;45(3):194-205; Mahnke, Clin Lab Med. 2007 Sep;27(3):469-85

## Slide 41
Spillover spreading across all PMTs
‘Nguyen, Cytometry A, 2013 Mar;83(3):306-15
Critical combinations
No spread areas

## Slide 42
Choose the best combo
HLA-DR Cy7APC
Concentration
Fluorescence
HLA-DR FITC
Concentration
Poor separation – Switch to Cy7APC
B cells
Activated T
Resting

## Slide 43
Panel-specific adjustments
CD95 Cy5PE
CCR7 BV711
Poor CD95 separation, but always looked fine in regular surface staining
TF buffer is harsh and may alter epitopes
Titration of CD95 with TF buffer
Repeat with adjusted Ab concentration
CD95 Cy5PE
CCR7 BV711

## Slide 44
Sample-specific adjustments
CD4 BV570
CD8 BV780
Blood
Lung tumor
“Optimal” CD8 titer according to titration in blood  too much spread onto other channels due to increased CD8 fluorescence in the tumor
Suboptimal CD8 titer to reduce data spread and increase sensitivity in tumor sample

## Slide 45
Sample-specific adjustments
CD45 Cy7APC
Aqua Live/dead
CD3 BV655
Dispensable CD45 staining, CD3 separates very well to identify T cells
Free Cy7APC channel to accommodate HLA-DR and detect activated cells
Lung tissue

## Slide 46
Fluor Antigen Notes
FITC HLA-DR
PerCP5.5 CD45RO
PE IRF4
PE-TxR Eomes
PE-Cy5 CD95
PE-Cy7 T-bet
APC CD45
Alexa700 GrB
Cy7APC CD69
BV421 PD-1
Dump Aqua
BV570 CD4
BV605 CD57
BV650 CD3
BV711 CCR7
BV780 CD8

## Slide 47
Fluor Antigen Notes
FITC HLA-DR
PerCP5.5 CD45RO
PE IRF4
PE-TxR Eomes
PE-Cy5 CD95
PE-Cy7 T-bet
APC CD45
Alexa700 GrB
Cy7APC CD69
BV421 PD-1
Dump Aqua
BV570 CD4
BV605 CD57
BV650 CD3
BV711 CCR7
BV780 CD8

## Slide 48
Fluor Antigen Notes
FITC HLA-DR Poor, move to Cy7APC
PerCP5.5 CD45RO
PE IRF4
PE-TxR Eomes Not that bright
Sensitive to CD57 spread
PE-Cy5 CD95 TF buffer sensitive
PE-Cy7 T-bet
APC CD45 Dispensable, drop
Alexa700 GrB
Cy7APC CD69 Fine, but move to FITC
BV421 PD-1
Dump Aqua
BV570 CD4
BV605 CD57 Very bright, titer down
BV650 CD3
BV711 CCR7
BV780 CD8 Titer down

## Slide 49
Fluor Antigen Notes
FITC HLA-DR Poor, move to Cy7APC
PerCP5.5 CD45RO
PE IRF4
PE-TxR Eomes Not that bright
Sensitive to CD57 spread
PE-Cy5 CD95 TF buffer sensitive
PE-Cy7 T-bet
APC CD45 Dispensable, drop
Alexa700 GrB
Cy7APC CD69 Fine, but move to FITC
BV421 PD-1
Dump Aqua
BV570 CD4
BV605 CD57 Very bright, titer down
BV650 CD3
BV711 CCR7
BV780 CD8 Titer down
Operations
Panel design
Titrate all reagents and predict spreading with titration curves
Test panel
Swap HLA-DR with CD69
Titrate new reagents
Test panel
CD57 still too bright, titer down, test with FMO controls
TF sensitive, re-titer CD95
Test panel on PB and tumor, drop CD45
Replace with Perforin, titrate Ab
EXPERIMENTS!!
1 month to develop 16 color panel, considering backbone

## Slide 50
Fluor Antigen Notes
FITC HLA-DR Poor, move to Cy7APC
PerCP5.5 CD45RO
PE IRF4
PE-TxR Eomes Not that bright
Sensitive to CD57 spread
PE-Cy5 CD95 TF buffer sensitive
PE-Cy7 T-bet
APC CD45 Dispensable, drop
Alexa700 GrB
Cy7APC CD69 Fine, but move to FITC
BV421 PD-1
Dump Aqua
BV570 CD4
BV605 CD57 Very bright, titer down
BV650 CD3
BV711 CCR7
BV780 CD8 Titer down
Final panel
CD69
CD45RO
IRF4
Eomes
CD95
T-bet
Perforin
GrB
HLA-DR
PD-1
Aqua
CD4
CD57
CD3
CCR7
CD8

## Slide 51
20-parameter FACS
FSC
SSC
BV421
BV510
BV570
BV605
BV650
BV711
BV786
FITC
Cy5.5PerCP
APC
Alexa700
Cy7APC
PE
PE-CF594
Cy5PE
Cy5.5PE
Cy7PE

## Slide 52
Moving towards 50 parameter FACS
BD Symphony A5
5 lsr: 395 405 488 561 640
BUV395
BUV496
BUV563
BUV615
BUV661
BUV737
BUV805
BB630
BB660
BB700
BB790
BV480
…and new IR and deepUV dyes

## Slide 53
Chart: 
| BB610 | PE-Cy5 | PE-CF594 | BB700 | BYG584 | BB670 | PE | BB515 | BB780 | BUV661 | BV750P | PE-Cy7 | BV421 | BV711 | BV480 | BUV805 | BV650 | BUV615 | BUV563 | BV605 | PE-Cy5.5 | BUV737 | APC-R700 | BV510 | APC-Cy7 | BUV496 | PCP-Cy5.5 | Alexa 488 | BUV395 | BV786 | Alexa 647 | BV570 | FITC
Stain Index Chart | 772.2182 | 766.6047 | 639.1231 | 608.1421 | 540.7091 | 537.7465 | 501.5 | 481.7207 | 422.7864 | 415.3958 | 339.075 | 266.9759 | 237.8387 | 230.175 | 217.6931 | 193.2295 | 187.875 | 179.6267 | 162.2706 | 158.4054 | 129.6909 | 112.1053 | 104.3056 | 83.1573 | 76.5538 | 73.2947 | 70.5538 | 66.3077 | 65.7273 | 60.4953 | 53.7963 | 51.1579 | 30.1944
Relative brightness
A new Age of Enlightenment in FACS

## Slide 54
Flow cytometry data presentation

## Slide 55
A call for data integrity in FACS
Science is experiencing an increased number of paper retractions
Frauds: can’t do anything unless giving better supervision
Data quality: gels, statistics, bioinformatics
Flow cytometry is widely used in immunology and hematology
Sample preparation is SIMPLE and for somewhat reasons people think they don’t need training (or need little)

## Slide 56
Introducing artifacts in flow cytometry
Instrument setup and calibration (PMT voltage generally dependent on the mood of the operator)
Reagent titration: high background, little separation
COMPENSATION
CONTROLS (define positivity)

## Slide 57
Some papers rely heavily on FACS

## Slide 58
SHOW ME THE DATA !
Some papers rely heavily on FACS

## Slide 59
Questions arising
Were the data compensated correctly?
And the gating? Same gates used for all patients/controls?
Were antibodies titrated?
How many cells were acquired? Treg cells are ~5% of CD4+, transplanted people are lymphopenic. Methods can report 1,000,000 cells were acquired – but what if the frequency of total CD3+ is 1%?
Cannot tell if raw data are not provided

## Slide 60
Outline
Data display methods
Dot vs. density plot vs. contour
Gating strategy
Type of gating

## Slide 61
Avoid histograms (or be careful)
Histograms are very intuitive but display only one parameter (we now live in a multicolor world)
88.1
77.8
CCR7
CCR7
Under
compensated
Correctly
compensated
CD4

## Slide 62
Dot vs. density plot vs. contour
Multiple ways to display data – use depends on applications
CD45RA
CCR7
Color
Contour
Density
Dot

## Slide 63
Saturated areas by events
CD45RA
CCR7
50,000
150,000
300,000
Number of events
300,000

## Slide 64
Dots are useful with low numbers
PBMCs were stimulated with pool of peptides to induce cytokine production (identification of ag-specific T cells)
Determine the memory phenotype of cytokine producing cells
IL-2
IFN-g
Unstimulated
Peptide stimulated

## Slide 65
Dots are useful with low numbers
CCR7
High resolution
(can’t see anything)
Low resolution
(better, but gates?)
No idea of number of events
CD45RA

## Slide 66
Dots are useful with low numbers
Dots overlaid on density plot
Contrasting colors
Total CD8+ cells allow to see population distribution
Total CD8+
Ag-specific

## Slide 67
Complex experiments: gating strategy
Why is this important? Tells you whether:
the authors have experience with FACS
gates were performed correctly (polygon vs. quadrants as a consequence of spreading)
Dead cells were excluded in the right way
There is contamination from unwanted populations
Cells were chosen correctly on the basis of FSC and SSC

## Slide 68
Memory T cells: full gating strategy
Is fluorescence stable over time?
Doublets excluded?
Dead cells excluded correctly? Contaminating populations such as Monos?
Only small lymphocytes or also large lymphocytes included?
Spreading of CD8 into CD4: optimal gate?

## Slide 69
MIFlowCyt – ISAC standards for data display

## Slide 71
enrico.lugli@humanitasresearch.it