-80 Freezer (Revco UXF) – Thermofisher

Serial Number: 0132864801140430

Model Number: UXF50086A

Item Number: 152AR5A01A

Service Team: AGR Mechanical

619-977-0271

agrmechanicalsd@gamil.com

Maintenance for the -80:

- Must be defrosted at minimum once a year

- The middle of the freezer must always be full

- If freezer is not completely full, fill it with empty shelves

- Clean the filter in the front every month

  - Slide out the metal guard

  - Slide out the filter

  - Wash with warm water and dry overnight

  - Replace filter and guard next morning

- Clean the filter on the bottom every month

  - Slide out the 2inch filter under the front side of the freezer

  - Wash with warm water and dry

  - Replace the filter that same day

- Shave off ice once a month

- Clean the compressors and system on the side of the -80 annually

  - AGR will do this as part of their service

- If it is not maintaining temperature or losing temperature quickly, call AGR Mechanical

  - They come out the same day or next day and will monitor the machine and give a diagnostics report

- Service the machine regularly

  - AGR has a service where they come out and clean the machine, check the system, monitor it for an hour, and regulate the compressors, etc

  - $135/service or $290/4 services in one year

*Note: Apparently our system, is the “older” version of this machine and has lower watts than the newer models, which is why it takes a little longer to get back to the correct temperature, but as told by the technician, this is expected with this model.

Last serviced by AGR Mechanical: March 16, 2021 (CT)