## Slide 1: Basic principles of flow cytometry
02/24/26

## Slide 2: How to avoid artifacts in flow cytometry
Reagent titration: how to achieve the best separation over background
Multicolor (>1 parameter) compensation
Spreading error in multicolor FACS
Presentation of high-quality data for publication

## Slide 3: How to avoid artifacts in flow cytometry
Reagent titration: how to achieve the best separation over background
Multicolor (>1 parameter) compensation
Spreading error in multicolor FACS
Presentation of high-quality data for publication

## Slide 4: Pro and cons of antibody titration
PRO
Can save a lot of money (100 tests can become 400, or more)
Better separation, less background
VERY important when performing multicolor analysis (can minimize ‘spreading’ in other channels)
CONS
None

## Slide 5: Abs: not just the manufacturer’s instructions
Titer:
Fresh PBMC stained with 7 two-fold serial dilutions of BV421 anti-CD14 (µL of Ab in 100 µL total volume for 20’ at RT)

## Slide 6: Abs: not just the manufacturer’s instructions
5
2.5
1.25
0.6
0.3
0.15
0.07
Titer
AIM: Choose the titer at which the best separation between the positive and the negative
is achieved without increase in the background

## Slide 7: Abs: not just the manufacturer’s instructions
5
2.5
1.25
0.6
0.3
0.15
0.07
Titer
✓
SNR: Signal-to-noise ratio
MFI: median fluorescence intensity
SNR: MFI (pos) / MFI (neg)
Titer
Chart: 
| 5 | 2.5 | 1.25 | 0.63 | 0.31 | 0.15 | 7.0000000000000007E-2
SNR | 219 | 230 | 247 | 283 | 265 | 264 | 249
SNR
✓
AIM: Choose the titer at which the best separation between the positive and the negative
is achieved without increase in the background

## Slide 8: Abs: not just the manufacturer’s instructions
Titer
Saturating titer: titer at which all receptors are saturated
(no increase in positive signal despite the increase of the Ab concentration)
Separating titer: titer at which separation between the positive and the negative is obtained

## Slide 9: What they say you can use… CD8α BV786 (53-6.7)
Thermo Fisher
BioLegend
BD
0.25 µg!
1.25 µL or 1:80
0.5 µg!
2.5 µL or 1:40
??? µg!

## Slide 10: What you can actually use…
0.016 µg!
0.078 µL or 1:1,250

## Slide 11
Critical parameters affecting titers
TIME: the longer the incubation the lower the titer needed to reach saturation
TEMPERATURE: some Abs stain better at certain temperatures, e.g. chemokine receptors give a much higher fluorescence when stained at 37°C compared to RT or 4°C
BUFFER: intracellular staining (Fix/Perm buffers) can alter the optimized titer  titrate antibodies with the buffer used in the staining
Cell number DOES NOT affect the Ab titer. The Ab is always in vast excess to the number of molecules, unless cell number changes dramatically.
General rule: the optimized titer can be used to stain up to 10-15x106 cells in 100 µL total volume.

## Slide 12: How to avoid artifacts in flow cytometry
Reagent titration: how to achieve the best separation over background
Multicolor (>1 parameter) compensation
Spreading error in multicolor FACS
Presentation of high-quality data for publication

## Slide 13
Compensation practices
Compensation can be performed manually (at the machine) or automatically (via software)
Manual compensation: tedious, need to check carefully that positive and negative signals have to match
Automated compensation: no errors, if using the right controls
Manual compensation is impractical in a multicolor experiment (roughly 2n pairwise combinations)
Changing compensation in a pairwise combination affects distribution in other channels

## Slide 14
What are the right controls?
The 3 golden rules of compensation
Use a single-stained control that is at least as bright as the test staining
Use the same reagent that is being used in the staining (valid for tandems such as PE-Cy7, APC-Cy7, etc. – for non tandems any reagent is OK)
Have a paired-negative control with the same autofluorescence as the positive (cells with cells, beads with beads, etc.)
Everything else is not necessary!

## Slide 15: Rule #1: at least as bright!
FITC IFNγ
Alexa 594 CCR7
Overcompensation of FITC into Alx594
Comp control
Test staining: IFNγ
Comp control is not as bright as the test stain

## Slide 16: Rule #2: same reagent for tandems!
Wrong comp
13.2% subtraction
PE-Cy5.5
Correct comp
3.4% subtraction
PE
Negative control
PE-Cy5.5
PE
Wrong
Correct
Compensation requirements can be different from lot to lot of the same Ab

## Slide 17: How to compensate in FACSDiva
Choose Experiment  Compensation setup  Create compensation controls
Run single-stained controls, including negative
Select positive and negative signal from each control
For comp tubes containing both positive and negative,
draw an additional gate to select the negative population
Choose Experiment  Compensation Setup  Calculate Compensation.
Enter a compensation setup name, click Link and Save, and click OK.

## Slide 18: How to compensate in FlowJo
Saves time at the instrument!
Create compensation controls as explained in the slide before
Import all files to FlowJo
1
2
3

## Slide 19: How to compensate in FlowJo
4
5
6
7

## Slide 20: Changing compensation affects distribution in other channels
FITC (em. 525) and PerCP
(em. 670) are distant enough
Possibly wrong comp between PerCP
(em. 670) and PE-TxR (em. 610)
START FROM SPLEEN (easy, standard, familiar)
THEN CHECK ALL TISSUES! And always keep your eyes open for artifacts
8
9

## Slide 21: How to avoid artifacts in flow cytometry
Reagent titration: how to achieve the best separation over background
Multicolor (>1 parameter) compensation
Spreading error in multicolor FACS
Presentation of high-quality data for publication

## Slide 22
‘Errors’ in compensation
Properly compensated data may not appear rectilinear because of measurement errors (are more prominent in case of limited numbers of photons, like PE-Cy7 or APC-Cy7 dyes, with low energy)
Measurement errors are directly proportional to the intensity (LOG relationship)
This effect on compensated data is unavoidable, and cannot be corrected with compensation (LINEAR correction)
Compensation does not introduce the error, it only reveals it because it brings the distribution to the low end of the log scale

## Slide 23: Signal spreading
Compensated
Compensated with transformation
Uncompensated
1000
2000
-500
+500
What causes this spread? Photon counting statistics introduces a measurement error

## Slide 24: How do we overcome this problem?
Use different gates: cells positive in the primary detector will have a different threshold of positivity compared to negative cells
Not a problem if analyzing bright antigens in the secondary detector
Loss of sensitivity if dim antigen in the secondary detector. Decrease titer of spreading Ab to minimize spreading.
-> Try a different Ab combination
APC
Alexa 700

## Slide 25: How do we overcome this problem?
APC
Alexa 700
100% detected
Use different gates: cells positive in the primary detector will have a different threshold of positivity compared to negative cells
Not a problem if analyzing bright antigens in the secondary detector
Loss of sensitivity if dim antigen in the secondary detector. Decrease titer of spreading Ab to minimize spreading.
-> Try a different Ab combination

## Slide 26: How do we overcome this problem?
Use different gates: cells positive in the primary detector will have a different threshold of positivity compared to negative cells
Not a problem if analyzing bright antigens in the secondary detector
Loss of sensitivity if dim antigen in the secondary detector. Decrease titer of spreading Ab to minimize spreading.
-> Try a different Ab combination
APC
Alexa 700
100%
detected
~50%
detected

## Slide 27: Fluorochrome relative brightness chart
Chart: 
| BB610 | PE-Cy5 | PE-CF594 | BB700 | BYG584 | BB670 | PE | BB515 | BB780 | BUV661 | BV750P | PE-Cy7 | BV421 | BV711 | BV480 | BUV805 | BV650 | BUV615 | BUV563 | BV605 | PE-Cy5.5 | BUV737 | APC-R700 | BV510 | APC-Cy7 | BUV496 | PCP-Cy5.5 | Alexa 488 | BUV395 | BV786 | Alexa 647 | BV570 | FITC
Stain Index Chart | 772.2182 | 766.6047 | 639.1231 | 608.1421 | 540.7091 | 537.7465 | 501.5 | 481.7207 | 422.7864 | 415.3958 | 339.075 | 266.9759 | 237.8387 | 230.175 | 217.6931 | 193.2295 | 187.875 | 179.6267 | 162.2706 | 158.4054 | 129.6909 | 112.1053 | 104.3056 | 83.1573 | 76.5538 | 73.2947 | 70.5538 | 66.3077 | 65.7273 | 60.4953 | 53.7963 | 51.1579 | 30.1944
Relative brightness

## Slide 28: How to avoid artifacts in flow cytometry
Reagent titration: how to achieve the best separation over background
Multicolor (>1 parameter) compensation
Spreading error in multicolor FACS
Presentation of high-quality data for publication

## Slide 29: Some papers rely heavily on FACS

## Slide 30: Some papers rely heavily on FACS
WHERE ARE THE DATA?

## Slide 31: Questions
Were the data compensated correctly?
And the gating? Same gates used for all patients/controls?
Were antibodies titrated?
How many events were acquired?
Cannot tell if raw data are not provided

## Slide 32: Careful with histograms
88.1
77.8
CCR7
CCR7
Under
compensated
Correctly
compensated
CD4
Histograms are very intuitive but display only one parameter
(we now live in a multicolor world)

## Slide 33: Dot vs. density plot vs. contour
Multiple ways to display data – use depends on applications
CD45RA
CCR7
Color
Contour
Density
Dot

## Slide 34: Dots are useful with low numbers
CCR7
High resolution
(can’t see anything)
Low resolution
(better, but gates?)
No idea of number of events
CD45RA

## Slide 35: Dots are useful with low numbers
Dots overlaid on density plot
Contrasting colors
Total CD8+ cells allow to see population distribution
Total CD8+
Ag-specific

## Slide 36: CD8 gating strategy
Is fluorescence stable over time?
Doublets excluded?
Dead cells excluded correctly? Contaminating populations such as Monos?
Only small lymphocytes or also large lymphocytes included?
Spreading of CD8 into CD4: optimal gate?

## Slide 37: How to avoid artifacts in flow cytometry
Reagent titration: how to achieve the best separation over background
Multicolor (>1 parameter) compensation
Spreading error in multicolor FACS
Presentation of high-quality data for publication

## Slide 38: Suggested read
https://www.nature.com/articles/s41596-019-0166-2